## Slide 1: DESIGN AND FABRICATION OF MJTCs ON QUARTZ SUBSTRATES AT NIST
L. Scarioni , T. E. Lipe and J. R. Kinard
Universidad de Carabobo Facultad Experimental de Ciencias y Tecnología
Valencia Venezuela
Notes: <number>

## Slide 2: NIST – SANDIA NATIONAL LABORATORIES THIN-FILM MULTIJUNCTION THERMAL CONVERTER
Silicon Frame
Thermocouples
Heater
Membrane
Notes: <number>
The standard PTB-PJMTC is fabricated on a silicon chip with a thin dielectric membrane. A thin-film bifilar heater and a up to 100 thermocouples are sputtered on an Si3N4/SiO2/Si3N4 sandwich membrane. The hot junctions of the thermocouples are located along the bifilar heater and the cold junctions are on he silicon frame acting as a heat sink. The window is anisotropically etched into the silicon chip.

## Slide 3: FABRICATION OF THE QUARTZ-MJTCs
Micromachining of Quartz
At-cut monocrystalline quartz wafer.
76 mm diameter, 250 mm thickness.
Quartz etchant: mixture of hydrofluoric acid, HF, and ammonium fluoride, NH4F in proportions HF:NH4F = 2:3. The etch rate is 20 m/h for a temperature of 60 °C.
Au layer of 100 nm thickness evaporated on a 10 nm Cr layer as mask material for etching. Etching patterns for the membrane were defined by standard photolitographic techniques.
Quartz Wafer
Quartz Membrane
Notes: <number>
For the construction of the new thermal-converter, we have used the AT-cut monocrystalline, because with this cut we obteined a smooth surface.
The wafer has a diameter of 76 mm and a thickness of 250 micrometers.
The quartz etchant was a mixture of hydrofluoric acid, HF and amonium fluoride, NH4F in proportions HF:NH4F = 2:3. The etch rate for this concentration is around 20 m/h for a temperature of 60°C.
An Au layer of 100 nm thickness evaporated on a 10 nm Cr layer was used as mask material for etching. The etching patterns for the membrane were defined in the mask by standard photolitographic techniques and carefully aligned at direction rotated by 24° against the crystallographic x-axis.
Thi figue show that the etching along the crystal axis results in a reproucible geometry of the window with a quartz membrane of a few micrometers thickness.

## Slide 4
Membrane Fabrication
Etching Beakers
Wafer Showing Membranes

## Slide 5: COAXIAL MJTC ON A QUARTZ CRYSTAL CHIP
Heater
Thermocouples
Membrane
Boundary
Ceramic
Mount
Die is 6 mm by 7 mm.
Notes: <number>
The structure of the bifilar heater and the thermocouples was fabricated using the standard photolithographic process and wet etching. The figure shows a photo of the thin-film MJTC chip on a quartz crystal chip.

## Slide 6
BIFILAR MJTC ON A QUARTZ CRYSTAL-CHIP
Heater
Thermocouples
Ceramic
Mount
Die is 6 mm by 7 mm.
Membrane
Boundary

## Slide 7: THE SIMULATION MODEL
Change of the real part of the heater impedance Z with frequency
Notes: <number>
To calculate the ac-dc transfer differences of the new thermal converter a model of the PJMTC was developed. This model includes the contribution of ac-dc differences arising from:
The distributed model considers the heater inductance, capacitive coupling between the two halves of the heater, capacitive coupling between the heater and the thermocouples and the power losses in the quartz substrate. The PMJTC heater is modeled as a lossy transmission line of ten elemens. The model is built from lumped elements i.e. resistances, capacitances, inductances and conductances.

## Slide 8
1
1
0
1
0
0
F
r
e
q
u
e
n
c
y
i
n
M
H
z
R
=
2
5
0

R
=
5
0
0

R
=
1
0
0
0

CALCULATED VALUES FOR DIFFERENT HEATER RESISTANCES
0
-500
-1000
-1500
-2000
Ac-dc Difference in µV/V

## Slide 9: HOUSING WITH INTEGRATED Tee FOR THE CALCULABLE QUARTZ-MJTC
Signal
Input
Defined
Signal
Output
Thermo-
Couple
Output

## Slide 10
1
1
0
1
0
0
0
Ac-dc differences in %
F
r
e
q
u
e
n
c
y
i
n
M
H
z
S
i
m
u
l
a
t
i
o
n
M
e
a
s
u
r
e
m
e
n
t
s
MEASURED AC-DC DIFFERENCES FOR COAXIAL QUARTZ MJTC
-1
-2
-3
-4

## Slide 11
1
1
0
1
0
0
Ac-dc differences in %
F
r
e
q
u
e
n
c
y
i
n
M
H
z
S
i
m
u
l
a
t
i
o
n
M
e
a
s
u
r
e
m
e
n
t
s
MEASURED AC-DC DIFFERENCES FOR BIFILAR QUARTZ MJTC
0
-1
-2
-3
-4
-5
-6

## Slide 12: PRESENT ACTIVITIES
New set of converters, with smaller heater resistances (≈ 830 Ω) is being fabricated.
Thicker thermocouple depositions are also being used, in order to reduce the thermocouple resistances from about 40 kΩ to less than 20 kΩ.

## Slide 13: CONCLUSION
A new generation of MJTCs on crystal quartz chips and membranes are being fabricated at NIST.
The results of the simulation show that the use of crystalline quartz allows significant reductions in the standard uncertainties and the ac-dc differences of the MJTCs in the frequency range from 100 kHz to 100 MHz.
Notes: <number>

## Slide 14: ACKNOWLEDGMENT
Research performed in part at the NIST Center for Nanoscale Science and Technology.
Funded as part of CCG
Project Number 546.

## Slide 15: Thank you for your attention
L. Scarioni , T. E. Lipe and J. R. Kinard