## Slide 1: RAOS Objectives
Characterize new and existing instruments for measuring the components of aerosol light extinction, with an emphasis on absorption
Quantify the uncertainty in the (filter-based) measurements of aerosol light absorption coefficient used by NOAA and ARM
Derive methods for determining spectral aerosol absorption from multi-wavelength measurements of absorption and (extinction - scattering)
Note: All experiments done at low RH, primarily external mixtures of black/white aerosols. Less complicated than atmospheric aerosols, perhaps not representative of them.

## Slide 2: Agreement of Absorption Standards
Notes: Data points represent kerosene soot runs with varying amounts of ammonium sulfate. Extinction data were adjusted to agree with scattering measurements on white aerosols (~5% adjustment to Eadj data).

## Slide 3: Reno2002 Light Absorption Comparison
Notes: Data points represent all low-extinction kerosene soot runs with varying amounts of ammonium sulfate. Bond et al. (1999) corrections applied to PSAP data. PSAP wavelength adjusted to 532 nm assuming a l-1 dependence.

## Slide 4: Summary of Results
Agreement of standard methods for absorption is ±5% at best
PSAP agrees with standard methods within about ±5% for typical atmospheric absorption levels
Filter transmission methods (PSAP, aethalometer) require improved corrections for multiple scattering effects
Cavity ringdown extinction cells appear to be biased low by 10-15% (relative to classic extinction and scat+abs)
All results are appropriate for low RH conditions (<20%), externally mixed black/white aerosols

## Slide 5: Continuing analyses
Compare results from kerosene soot with results from diesel soot and carbon vane pump exhaust
Compare run-averaged results to time-series data (1-minute averages)
Derive wavelength-dependence of absorption, to allow application of extinction-minus-scattering calibrations to other wavelengths
Derive correction schemes for filter-based methods (PSAP and aethalometer) at all measured wavelengths
Investigate reasons for the lower CRD extinction
Compare results from ambient aerosol samples