## Slide 1: NIST Training Seminar Augusta Maine August 14-18, 2000
Loading Rack Meters

## Slide 2: Loading Rack Meters
EPO No. 25
  Safety
  Pretest Determinations
  Inspection
  Test Notes
  Test Procedure

## Slide 3: Safety

## Slide 4: Safety
Clothing
  non-Synthetic
Electrical Hazards
  Exposed wiring
  Sources of Ignition (heat or spark)
Emergency Procedures
  Phone Numbers for Emergency Services
  Basic First Aid Training
Eye Protection
  Goggles or Safety Glasses
  Eye Wash

## Slide 5: Safety Cont.
Fire Extinguisher
  Portable (20 lb or larger)
  Dry Chemical
First Aid Kit
  Basic Supplies
Grounding
  Grounding Cable and lugs
  Skully System

## Slide 6: Safety Cont.
Ignition Sources
  Open wiring
  Exposed Lights
  Improper Extension Cords
  Portable Heaters
Lifting
  Proper Lifting Techniques
  Weight Limitations
Location
  Condition (wet/slick )
  Obstructions, Overhead Hazards
  Petroleum Products, Hazardous Materials of Chemicals (MSDS’s)

## Slide 7: Safety Cont.
Material Data Sheets
Nature of Product
Personal Protection Equipment
  Safety Shoes
  Protective Clothing, Gloves, Eye Wear, Respirator, Barrier Cream
  Hard Hat
Safety Cones/Warning Signs
Static Discharge
  Grounding
  Clothing

## Slide 8: Safety Cont.
Support for Prover
  Proper Leveling Jacks
  Wheel Chocks
Switch Loading
  Diesel Fuel First
Traffic
Transportation of Equipment

## Slide 9: Pretest Determinations

## Slide 10: Pretest Determinations
Prover
  Must have a current calibration certificate
  Seals must be in place
  Design compatible with system (top/bottom loading
  Capacity sufficient to hold one minute flow at max rate and not less than 200 L ( 50 gal ) N.3.5.
  Condition - no dents, rust, water, or foreign material , sight glass clean and not leaking
  For top loading prover inlet lower than meter outlet
Thermometers
  Range 0 0F to 120 0F , 0.5 0F increments, 0.2 0F accuracy

## Slide 11: Pretest Determinations Cont.
Test Fluid N.1.1.
  Same characteristics as that measured commercially
Applicable Requirements G-T, T.1.
  Basic Tolerances T.2.3.3.
    Normal tests
      Acceptance tolerance 0.2 %
      Maintenance tolerance 0.3 %
    Special tests
      Acceptance tolerance and
      Maintenance tolerance 0.5 %

## Slide 12: Inspection

## Slide 13: Inspection
Indicating and recording elements
  Design
    Must have indicating element, may have recording element S.1.1.
    Units in liters or gallons or binary-submultiples or decimal subdivisions S1.2., S.1.2.3.
  Readability
    Indicating and recording elements clear, definite, and easily read G-S.5., S.1.4., S.1.5.
    Required markings distinct, easily readable and permanent G-S.6., G-S.7.

## Slide 14: Inspection
Indicating and recording elements cont.
  Value of intervals
    Uniform throughout series G-S.5.3.
    If designed to indicate more than one unit - value must be appropriately identified G-S.5.3.1.
  Advancement and return to zero
    Elements may only be advanced if S.1.3.
      Cannot be stopped until zero is reached or
      Indications are automatically obscured until zero is reached

## Slide 15: Inspection
Indicating and recording elements cont.
  Provision for sealing
    Provision must be made for sealing electronic adjustable components G-S.8.
    A security seal must be affixed to any adjustment mechanism designed to be sealed G-UR.4.5.

## Slide 16: Inspection
Measuring elements
  Effective vapor elimination system S.2.1.
    Vent lines suitably rigid
  Means provided for determination of product temperature S.2.6.
  Provision for applying security seals to meter and automatic temperature compensating system
    All security seals intact G-UR.4.5., S.2.2., S2.7.3.

## Slide 17: Inspection
Installation
  Readily accessible for testing G-UR.2.3., G-UR.4.4.
    Assistance provided if required
  Liquid cannot be diverted from measuring chamber or discharge line and directional flow valve are automatic S.2.3., S.3.1.
  No leaks on meter outlet side G-UR.4.1., S.3.1.
    Leaks on inlet side noted and called to users attention

## Slide 18: Inspection
Installation cont.
  System and associated equipment installation, and construction do not facilitate fraud G-S.2.
  Installation is proper and does not adversely affect system performance G-UR2.1., UR.2.1.
  Actual maximum discharge rate does not exceed manufacturer’s specification UR.2.2.

## Slide 19: Inspection
Selection and use
  Device suitable for application G-UR.1.1., G-UR.1.2.
    Design
    Flow rate
    Computing capability
    Indicating and recording elements
    Value of smallest unit and unit price
    Environment

## Slide 20: Inspection
Selection and use cont.
  Device and associated equipment operated in manner intended by manufacturer G-UR.3.1., G-UR.4.1.
  If equipped with mechanical ATCS UR.3.6.1.1.
    Connected and operable
    In use at all times

## Slide 21: Inspection
Marking
  Make, model, and serial number G-S.1.
  Switches, lights, displays, pushbuttons, controls, and features clearly identified G-S.6.
  Limitations of use clearly and permanently marked S.4.1.
  Minimum and maximum discharge rate clearly marked S.4.3.1.
    Minimum cannot exceed 20 % of maximum

## Slide 22: Inspection
Devices equipped with ATCS
  Provision to deactivate S.2.7.2.
  Thermometer well provided for determination of product temperature S.2.7.4.
  Primary indicating and recording elements marked to show delivered volume has been adjusted to 60 0F S.4.3.2.

## Slide 23: Test Notes

## Slide 24: Test Notes
Level Prover
  Level indicators working properly
Connect safety interlock, ground cable and vapor recovery hose if applicable
Verify all valves closed and pumping mechanism functioning
Record totalizer reading
  Before and after each draft

## Slide 25: Test Notes
Minimize temperature differences between meter and prover N.2.
Minimize splashing in top loading provers and maintain fill spout in consistent position
Examine printed tickets and invoices
  Print after every draft G.S.5.6.
  Computing type
    Check computations S.1.7.2.
    Total volume and unit price must accompany total price either printed or handscript UR.3.4.

## Slide 26: Test Notes
Examine printed tickets and invoices cont.
  Check indicated and recorded values for comparability G-S.5.2.2.
  With ATCS check invoices for UR.3.6.1.2. a & b
    Notice that volume has been adjusted to 60 0F
    For electronic systems, listing of API gravity, specific gravity, or coefficient of expansion, product temperature, and gross meter reading
  Without ATCS UR.3.6.2.1., UR.3.6.2.2.
    If adjusted, notice that volume has been adjusted to 60 0F
    Temperature of product used to make adjustment

## Slide 27: Test Notes
Reading meniscus
  Bottom for transparent liquids
  Top for opaque liquids
Draining prover
  Lower sight glass
    Close drain valve before liquid reaches zero, wait 30 seconds, open small zero drain valve till zero
      do not readjust
  No lower sight glass
    When drips start wait 30 seconds, close drain

## Slide 28: Test Notes
Temperature readings
  Read to nearest 0.5 0F
    Use matching thermometers
  Read prover temperature immediately after each accuracy test run
    Average readings if prover has multiple wells
  Read meter temperature
    Thermometer well in measuring chamber or immediately adjacent
    Read at 1/3 and 2/3 of prover capacity

## Slide 29: Test Notes
Automatic stop mechanism (preset) must stop within one-half the minimum indicated interval G-UR.4.1.

## Slide 30: Test Procedure

## Slide 31: Test Procedure
Wet prover
Empty prover using 30 second drain
Insert ticket & set preset to prover volume
Zero meter
Start pump then open delivery valve
At end of each run read prover & determine meter error
    Repeat run if test result is near or out of tolerance

## Slide 32: Test Procedure
Notes for Repair Personnel
  At least two runs at same flow rate to determine repeatability
    Individual runs within applicable tolerance
    Difference not to exceed 40 % of applied tolerance
  Adjust meter if out of tolerance
    Deactivate ATCS if necessary
    Check performance of temperature sensor

## Slide 33: Test Procedure
Accuracy tests - no ATCS N.5.
  Normal test - full flow N.4.1., N.4.1.1., T.2.3.
    Flow rate within meter rating G-UR.3.1., G-UR.2.2.
  Special test - slow flow N.4.2., N.4.2.4., T.2.3.
    Fill prover at minimum flow rate marked on meter
Read temperature at meter and prover make corrections for meter and prover readings

## Slide 34: Test Procedure
Accuracy tests - no ATCS N.5.
  Normal test - full flow N.4.1., N.4.1.1., T.2.3.
    Flow rate within meter rating G-UR.3.1., G-UR.2.2.
    Average product temperature at meter - read at 1/3 & 2/3 of prover capacity
    Product temperature in prover
Correct prover, product in prover, & meter readings for temperature difference from 60 0F

## Slide 35: Test Procedure
Accuracy tests - no ATCS N.5.
  Special test - slow flow N.4.2., N.4.2.4., T.2.3.
    Fill prover at minimum flow rate marked on meter
    Average product temperature at meter - read at 1/3 & 2/3 of prover capacity
    Product temperature in prover
Correct prover, product in prover, & meter readings for temperature difference from 60 0F

## Slide 36: Test Procedure
Accuracy tests of temperature compensated meter - indication in “net” gallons
  Normal full flow test - With ATCS activated
    Average product temperature at meter - read at 1/3 & 2/3 of prover capacity
      Only to note large variation
    Product temperature in prover
  Correct prover reading & product in prover for temperature difference from 60 0F

## Slide 37: Test Procedure
Accuracy tests of temperature compensated meter - ATCS deactivated
  Normal full flow test
    Average product temperature at meter - read at 1/3 & 2/3 of prover capacity
    Product temperature in prover
  Correct prover, product in prover, & meter readings for temperature difference from 60 0F

## Slide 38: Test Procedure
Accuracy tests of temperature compensated meter
  Special test slow flow - ATCS deactivated
    Average product temperature at meter - read at 1/3 & 2/3 of prover capacity
    Product temperature in prover
  Correct prover, product in prover, & meter readings for temperature difference from 60 0F

## Slide 39: Test Procedure
Accuracy tests - For meters that display both “net” and “gross” deliveries
  For “gross” delivery indication use correction procedure for meter without temperature compensation for normal & special tests
  For “net” delivery indication use procedure for meter with ATCS for normal & special tests

## Slide 40: Test Procedure
For system with ATCS
  Difference in meter error with ATCS activated and ATCS deactivated shall not exceed
    0.2 % for mechanical ATCS and
    0.1% for electronic ATCS T.2.3.3.

## Slide 41: Test Procedure
Security seal
  Provision for security seals S.2.2.
  Seals in place & intact G-UR.4.5.
Wet-hose system S.3.9.
  Check effectiveness of anti-drain valve
Dry-hose system S.3.6.
  Check for complete drainage of hose

## Slide 42: Test Procedure
RFI/EMI
  Test with associated equipment only
    Test only if a problem is suspected
    G-N.2., G-UR.1.2., G-UR.3.2., G.UR.4.2.